## Slide 1: Mass Flow Sensor Integration
Team 5:
Beau Rodgers
Luis Mendez
Brian Roberts
Keenan Cheeks
Sponsored by: Danfoss Turbocor
Advised by: Dr. Shih
Notes: Hi I‘m Keenan cheeks, this is brian Roberts and we are team 5 mass flow sensor integration. We are working together with our sponsor Danfoss Turbocor, and our faculty advisor Dr. shih.

## Slide 2: Danfoss Turbocor
Danfoss Turbocor is the world leader in oil-free centrifugal compressors used for cooling applications.
Danfoss Turbocor is currently only able to measure compressor efficiency on site in their test facility.
Probe style in-line mass flow sensor lends itself to extensive modification of existing systems.
Figure 1: Endress-Hauser flow meter
Presenter : Keenan Cheeks
Notes: To begin I want to give a little background on Turbocor and what they do. Danfoss Turbocor has designed maglev technology that uses magnetic bearings to levitate the compressor shaft eliminating the need for oil.
Currently Turbocor can only measure the efficiency of these compressors at their testing facility using a very bulky in-line mass flow sensors.

## Slide 3: Developing a Project Scope
Initial scope: Determine failure modes using data analysis by implementing algorithms.
First Revision: Integrate a mass storage device onto the Variable Twin Turbo “VTT” compressor.
Second Revision: Develop an externally mounted mass flow sensor to determine compressor efficiency.
Presenter : Keenan Cheeks
Notes: Our first major task was to determine what our sponsor Turbocor wanted from us that would be helpful to them in the future but also comply with the requirements of this class. We jumped from designing algorithms to determine failure modes to mass storage device integration and finally to what we are working on today with developing an externally mounted mass flow sensor package to determine compressor efficiency.

## Slide 4: Customer Needs
Need Statement: Turbocor seeks to add value to their existing product by determining real time efficiency.
Goal Statement: Integrate external mass flow sensor with the ability to measure real time efficiency of Turbocor compressors.
Presenter : Keenan Cheeks
Notes: Currently, more companies are starting to develop oil free compressors making the field more competitive. To stand out on top, Turbocor wants to develop a failure modes prediction system to provide to their customers. As of now Turbocor does not have a way of acquiring real time data or monitoring real time efficiency of their compressors.
Our goal here is to concentrate on providing a means by which Turbocor can measure real-time efficiency on existing Turbocor platforms that can easily be used with their current compressor monitoring system. This is a stepping stone in developing a failure modes prediction system.

## Slide 5: HOQ
Presenter : Keenan Cheeks
Figure 2 – House of Quality
Notes: Using a house of quality we determined that the most important requirement for our sponsor is that the inlet and exit temperature sensors be accurate. We can also see here that the mass flow sensor must be externally mounted so that existing customers won‘t have to motify their systems.

## Slide 6: Our Needs
Determine relevant thermodynamic properties and required parameters.
Develop relations/equations to calculate desired values.
Use Turbocor supplied software and data to test equations.
Determine best data collection sensors.
Analyze sensor output to determine compatibility.
Test system and determine errors and possible remedies.
Presenter : Keenan Cheeks

## Slide 7: System Thermodynamics
Typical Turbocor compressor applications are in water cooled HVAC systems.
Project focus will be on the evaporator portion of the HVAC system.
Presenter : Keenan Cheeks
Figure 3 : System Configuration
Figure 4: System Example
Figure 5 : Detailed Configuration
Notes: Turbocor compressors are used on water cooled HVAC system chillers to drive R-134a refrigerant. Here we can see a simplified picture of a water cooled HVAC system.
We have been advised to focus our efforts on the chiller and more specifically on the evaporator side of the chiller because of the difficulties that arise when dealing with superheated gas. Another reason we want to focus on this side of the chiller is because this is where the system cooling load is applied.
Here is a picture of a chiller with a Turbocor compressor. On top is the evaporator and on the bottom is the condenser. These pipes are the inlet and exit of the refrigerant and here are the inlet and exits of the water.
Here is a more detailed diagram of how the chiller functions with important parameters.

## Slide 8: Thermal-Fluid Correlations
Where:
mass flow rate of water
mass flow rate of refrigerant
= Inlet water temperature
exit water temperature = inlet enthalpy of refrigerant
= exit enthalpy of refrigerant
= inlet pressure of refrigerant
= exit pressure of refrigerant
= volume flow rate of refrigerant
Presenter : Keenan Cheeks
Notes: Here we have the heat transfer concepts that are involved with the evaporator
The first two equations are basic expressions for the rate of heat transfer experienced by a fluid. As seen in the previous figure the rate of heat transfer experienced by the water theoretically is the same heat transfer experienced by the refrigerant. This assumption allows us to find the rate of heat transfer of the refrigerant and ultimately the mass flow rate of the refrigerant without having to cut any pipes to insert an in-line mass flow sensor in the compressor cycle.
After we determine the mass flow rate of the refrigerant the next step is to calculate the power that the compressor is out putting by cooling the water using the third correlation.
Finally, the power output can be used to calculate the efficiency of the compressor.
The issue here is that we have made a major assumption that the evaporator transfers 100% of the heat or that it is 100% effective when we know that this isn’t true.

## Slide 9: Evaporator Effectiveness
Where:
= maximum possible heat transfer rate
= the smaller of
= effectiveness of heat exchanger
Presenter : Keenan Cheeks
Notes: After meeting with Dr. Kumar and discussing our approach to the problem he suggested that instead of making this assumption and introducing large errors in compressor efficiency calculations, we should calibrate the heat exchanger. This simply means finding the effectiveness of the heat exchanger. Furthermore it would be possible to find a complete calibration curve for the heat exchanger based off of the effectiveness NTU method.
Once the effectiveness of the heat exchanger is known the actual heat transfer rate can be found along with refrigerant mass flow rate.

## Slide 10: Necessary Parameters
Available Parameters
Required Parameters
Presenter : Keenan Cheeks
Figure 6 : Thermodynamic Cycle
Notes: From all of these correlations we currently have access to: available parameters
We will then need to measure: required parameters
And finally we can find work out and efficiency.

## Slide 11: Mass Flow Sensors
Available Technologies
    Coriolis flow measuring principle (in-line)
    Electromagnetic flow measuring principle (in-line)
    Thermal flow measuring principle (probe style)
    Vortex flow measuring principle (in-line)
    Ultrasonic Mass Flow Sensor (externally mounted)
Presenter : Brian Roberts
Notes: While researching the best way to calculate the mass flow rate without being intrusive to the current customer platforms we came across many different principles that can provide the necessary information. The only one that we found that was externally mounted was the ultrasonic mass flow sensor which was found to also be very accurate and robust.

## Slide 12: Design Concepts
Use existing ultrasonic transducers on the market for retrieving data and use developed algorithms and processing units to determine mass flow.
Use existing ultrasonic sensors with corresponding modules to determine mass flow rate and focus on integrating sensor outputs into existing Turbocor platforms.
Presenter : Brian Roberts
Notes: After choosing a direction and examining the properties of the system we had two basic approaches we came up with. We could use existing ultrasonic transducers and compute our own algorithms for analyzing the output. This would also entail designing the assembly with to mount the system on the chiller. Our other idea was to find a complete sensor set that we could interface with the existing compressor hardware. This would allow us to focus on integration and the mounting mechanism itself.

## Slide 13: Concept Evaluation
Not using existing data collection system would be cheaper but cause unnecessary work and complications.
Best choice would be to use existing sensors and corresponding electronics to retrieve mass flow data and focus on mounting apparatus and accurate compressor efficiency values.
Presenter : Brian Roberts
Notes: The obvious advantage to designing our own receiver for the transducers is cost. It would likely be cheaper to create the device ourselves than purchase it from a company. However the investment of time and extra level of complexity that adds to the project made it unfeasible. It was decided the best choice would be to pursue an existing sensor package and design modifications that would allow it to communicate with Turbocor’s hardware.

## Slide 14: Potential Challenges
Maintaining budget (~$2000)
Flow measurement accuracy (<1% error)
Clamp versatility
Compatibility of systems
      Sensors must have output that can be integrated into Turbocor’s existing platform.
Calibration of System
Presenter : Brian Roberts
Notes: The direction we chose has its own perils. First, these sensors tend to be expensive. Research showed systems can range from hundreds to thousands of dollars. Another concern is accuracy. Our sponsors want less than 1% error in our measurement. Another potential problem is compatibility. There are only a small number of unused inputs and our system must be able to interface with what is available. In addition it would be beneficial if the system needed as little calibration as possible for field setup. Part of the entire concept of automated data analytics is reducing the need for onsite visits from technicians. If the system is too difficult to setup it would be counter to this ideal.

## Slide 15: Sensor Selection
Ultrasonic Flow sensors:
    Doppler vs Transit Time
    Narrow Beam Ultrasonic
    Wide Beam Ultrasonic
Key Parameters
    Pipe
        Diameter
        Wall thickness
        Construction
    Configuration
    Fluid
Presenter : Brian Roberts
Figure 7 : Ultrasonic Sensor Fundamentals
Figure 8 : Different Configurations
Notes: The ultrasonic mass flow sensor is highly customizable depending on the customers needs. There are two main types of ultrasonic sensors, Doppler effect and time transit. The Doppler shift method is better for non-homogenous fluids. The time transit method tends to be more accurate and has two different technologies, wide beam and narrow beam. Both types are very accurate but the wide beam sensor allows for monitoring of a larger range of mass flow rates since it emits and receives a larger sound wave through the fluid. The wider beam can be detected even if the mass flow rate is increased by a large amount without sacrificing any performance. The sensors can come in different configurations as well. Shown here are the “Z” and “V” configurations.

## Slide 16: Sensor Selection Matrix
Characteristics Weighted 1-5 | Accuracy | Cost | Pipe ᴓ (m) | Fluid Temp(°C) | Fluid Velocity(m/s) | Calculate Q̇ ? | Output (Analog or Digital) | Total
 | 5 | 3 | 4 | 4 | 4 | 3 | 2 | 
 | Actual Values rated from 0-3 |  |  |  |  |  |  | 
Sierra InnovaSonic 210i | ±0.5% 3 | $3,170 1 | 0.15 to 6 3 | -40 to 80 2 | 0.05 to 12 1 | No 0 | A 2 | 46
Dalian Hipeak TDS-100 | > ±1% 2 | $500-700 3 | 0.15 to 7 3 | 0 to 90 2 | -30 to 30 3 | Yes 3 | A/D 3 | 66
GE Panametrics PT878 | ±1% 1 | $10,035 0 | .13 to 7.6 3 | -10 to 50 1 | -40 to 40 3 | No 0 | A/D 3 | 39
Greyline TTFM 1.0 | ±1% 1 | $2,600 2 | 0.12 to 1.2 2 | -40 to 150 3 | 0.02 to 12 2 | No 0 | A 2 | 43
Presenter : Brian Roberts
Table 1 : Selection Matrix
Notes: Several models of sensors were researched and analyzed. The best four out of the packages found are presented here in a selection matrix. The customer requirements were weighted from 1 to 5 and each model was assigned a value from 0 to 3 corresponding to how well it meets the design criteria. From this analysis it was determined that the Dalian Hipeak TDS-100 would best suit our needs.

## Slide 17: TDS-100
Large range of pipe diameters
  15 cm to 7 m
Appropriate range of fluid temperatures
  0-90°C
Adjusts for current fluid temperature
Mounting points included
Versatile Output
  RS-485
  Analog 4-24mA
Presenter : Brian Roberts
Figure 9 : TDS System
Figure 10 : TDS-100 Control Board
Notes: The TDS-100 is a versatile machine. Depending upon the configuration, it can measure velocities on pipes from 15cm to 18m. With added temperature sensors it can also adjust for the changing properties of the fluid due to differences in temperature in its operating range. Another benefit for our purpose is the included mounting points which can be used to integrate the TDS-100 into our mounting apparatus. In addition there are multiple types of output including the standard RS-485 communication protocol or an analog signal that ranges from 4-24mA. Both of these options would work for the existing system.

## Slide 18: Transducers
Different models for different diameters
Same control system can use multiple types of transducers
Mounted directly to outside of pipe
Presenter : Brian Roberts
Figure 11 : Basic Transducers
Figure 12 : Standard “V” Configuration
Notes: In order to measure the flow rate in different diameter pipes, different models of transducers need to be used. However the upside is that the same control unit can handle multiple types of sensors. The transducers used by these systems are mounted directly to the outside of the pipe as shown before.

## Slide 19: Temperature Sensors
Turbocor currently uses the TI 112CP3-4 sensor
  Detects pressure and temperature
  System already calibrated for use
Operating Ranges:
    -40 to 120°C
    9 to 149 PSI
Accuracy ±0.6%
Presenter : Brian Roberts
Figure 13 : TI 112CP3-4 Temperature and Pressure Sensors
Figure 14 : TI 112CP3-4 Connector Layout
Notes: Turbocor currently uses the Texas Instruments 112CP3-4 pressure and temperature sensor. Their equipment is designed to mount and read these sensors. In addition from our research it appears that the TDS-100 will be compatible with them as well. They are capable of exceeding the ranges being measured in the evaporator and they are accurate to 0.6%.

## Slide 20: Future Plans
Use existing experimental data to test validity of our thermodynamic model.
Order sensors and corresponding electronics.
Use hardware dimensions to begin CAD design of mounting clamp.
Collect and process new experimental data
    Calculate temperature differential
    Determine and use temperature and mass flow relationship to calculate efficiency of compressor
Presenter : Brian Roberts
Notes: Moving forward we need to begin acquiring the physical parts of our system and designing the mounting bracket. After we assemble the system we can start running tests to analyze our thermodynamic model. The next step will be to integrate the system and determine overall efficiency for the compressor.

## Slide 21: Project Plan
Figure 15 : Gantt Chart
Notes: Here is the Gantt chart we are currently using to guide our progress. Currently we are finalizing our needed parts and working on designing the bracket itself. We would like to get to prototyping early in the spring semester so that we can quickly move on to running experiments on the system.

## Slide 22: Conclusion
As new technologies begin to challenge Turbocor’s dominance, they desire added value for their products to maintain market share.
One way to achieve this goal is to add the ability to analyze data both instantly and over an extended period.
To aid in the goal of advanced data analytics we are designing a sensor module that can determine compressor efficiency to integrate into their platform.
Measuring efficiency is a necessary stepping stone towards advanced failure modes analysis.
Presenter : Brian Roberts
Notes: As any new technology become more common place, competition becomes ever fiercer for the originator of the technology. Turbocor is looking forward and trying to maintain their place at the top of the frictionless compressor market. One approach they are taking is to add features to their existing product, specifically data analytics. This allows for failure prediction before they occur and prevent unnecessary damage to equipment. To achieve their goals several different metrics need to be monitored. Compressor efficiency is one important aspect that can reveal system health. To help them in their goal we are working to determine the mass flow rate of the refrigerant which they are not currently able to monitor.

## Slide 23
Questions?

## Slide 24: References
https://www.gemeasurement.com/flow-measurement/ultrasonic-liquid/transport-pt878-portable-ultrasonic-liquid-flow-meter
      Ge Panametrics PT878
http://www.sierrainstruments.com/products/210Prod.html
      Sierra InnovaSonic 210i
http://www.ecefast.com.au/media/blfa_files/TDS-100Fmanual_ece.pdf
      Dalian Hipeak TDS-100F
http://www.greyline.com/ttfm10.htm
      Greyline TTFM 1.0
http://www.us.endress.com/en
      General Massflow Sensor information
http://www.cdiweb.com/Manufacturers/sensata/FP/Pressure_Transducer-Transmitter/?type=94&manf=512&cate=512:10&NavType=2#null
      TI 112CP3-4