## Slide 1: Apollo Renal Therapeutics Water Filtration Project
Team 504
Kencin Autry, Matthew Kennedy, Timothy Norman,
Lapadre Proctor, Lily Thompson

## Slide 2
Team Introductions
Kencin Autry
2
Kencin Autry
Research
Development
Engineer
Lapadre Proctor
Test and Quality
Control Engineer
Matthew Kennedy
Structural Design
And Materials
Engineer
Timothy Norman
Structural Design
And Manufacturing
Engineer
Lily Thompson
Computational
Fluids Engineer

## Slide 3
Sponsor and Advisors
Sponsor
Gary Mishkin
CEO Apollo Renal Therapeutics
DBA Artemis Plastics
Academic Advisor
Dr. Rebekah Sweat
Assistant Professor
Kencin Autry
Academic Advisor
Aspen Reyes Graduate Research Assistant

## Slide 4
Objective
The objective of this project is to reduce the quantity of waste-water in the dialysis process by efficiently purifying standardized filtered water and spent dialysate to ultra-pure water.
Kencin Autry
Notes: l

## Slide 5: What is Dialysate?
Kencin Autry

## Slide 6
6
Project Background
Hemodialysis is a therapy treatment for patients with end-stage kidney disease (ESKD).
Kencin Autry

## Slide 7
7
Hemodialysis Process
Kencin Autry

## Slide 8
8
Our System
RO System
Filtered
Water Runoff
Check Valves
Primary Pump
CNT Filter
Water Quality Testing
Water Quality Testing
Spent Dialysate Runoff
Kencin Autry

## Slide 9
9
Key Goals
Enhancing efficiency and preventing the potential hazards of large water filtration systems
Recycling spent dialysate back into ultra-pure water
Filtered water will meet ultra-pure water standards
Lapadre Proctor

## Slide 10
Lapadre Proctor
10
Customer Needs
10
10
Design requires relatively low maintenance and will last approximately 1 year
Design should fit into a space no larger than 10 in x 12 in x 5 in
Hansen connector and Barbed Port input
Handle a maximum flow rate of 1000 mL per minute
Produce Ultra-Pure water

## Slide 11: Critical Targets and Metrics
Purify Water
Recycle Spent Dialysate
Ergonomic Compatibility
Raise and Lower Flow Rate
< 0.001 ppm TDS
75% Efficiency
10 x 12 x 5 in
100 – 1000 mL per minute
Lapadre Proctor

## Slide 12: Concept Generation
Secondary Dialyzer
Lego Sorter
Nanofiltration and Reverse Osmosis
Chemical Coagulant
Nanofiltration
Lapadre Proctor

## Slide 13: High-Fidelity
Carbon Nanotube (CNT) and Ultrasonics
Forward Osmosis
Gaseous Ozone
Lapadre Proctor
Ozone Generator

## Slide 14: Final Concept
Reverse Osmosis (RO)
Carbon Nanotubes (CNTs) and Ultrasonics
Lapadre Proctor
Notes: * NASA

## Slide 15
RO System
Filtered
Water Runoff
Check Valves
Primary Pump
CNT Filter
Water Quality Testing
Water Quality Testing
Spent Dialysate Runoff
Lapadre Proctor
Notes: Needs to be cleaned up if we do use it

## Slide 16: RO System
3 Stage Filtering
2-in-1 Polypropylene Sediment and Carbon Block Filters
Specifications
Incoming water must be between 14.5 and 58 psi
Reverse Osmosis Membrane
Filters at 1590 mL per minute
Lapadre Proctor
Notes: Should we go into the individual components?
Claims to remove 99% of contaminants for drinking water standards
Filters last 12 – 24 months
Incoming water must be between 14.5 and 58 psi
Filters at 1590 mL per minute
Will have to adapt to two water inputs

## Slide 17
RO System
Filtered
Water Runoff
Check Valves
Primary Pump
CNT Filter
Water Quality Testing
Water Quality Testing
Spent Dialysate Runoff
Lily Thompson

## Slide 18: Primary Pump
Lily Thompson
Connections
Arduino Microcontroller Mega
Specifications
Maximum flow rate of 1600mL per minute
TB6600 Motor Driver
12 V Power Supply
Stepper Motor
Peristaltic
12 V Supply
Notes: Should we go into the individual components?
24V Stepper motor
Silicone tube,7.9mm ID × 11.1mm OD
Flow Rate of 1600ml/min
Ambient temperature 0-40℃
Relative humidity <80%
Reduction ratio 1:20

## Slide 19
RO System
Filtered
Water Runoff
Check Valves
Primary Pump
CNT Filter
Water Quality Testing
Water Quality Testing
Spent Dialysate Runoff
Timothy Norman

## Slide 20: Carbon Nanotube (CNT) Housing
Materials
Inside Layer: 3D Printed with PLA
Configuration
Well-rounded with angle of 45 degrees
Cylindrical
Middle Layer: Foam for ultrasonics
Outside Layer: PLA
Timothy Norman

## Slide 21: Ultrasonics
Timothy Norman

## Slide 22: Prevention
Lily Thompson
Notes: Ensuring the spent dialysate can’t leak around the CNT when placed into the housing.
FSU red flagging the Urea

## Slide 23: Configuration
Lily Thompson
Notes: Put dimensions on here (For this slide, How about inside of showing the angles on the exterior of the shell, we take a cross section of the lid and show the inside of it) - Matt

## Slide 24: Fluid Computations
Lily Thompson
Control Volume
Semi-permeable Membrane
Mass Conservation
?
Key
?
Notes: Put dimensions on here

## Slide 25: Take Away
Sudden Expansion:
Gradual expansion and contraction:
Future experiments can determine pressure difference, output velocity, and ideal ratio of diameters
Gradual expansion and contraction
Too many unknowns in calculations
Lily Thompson

## Slide 26: Inside Housing
Timothy Norman
Notes: I honesty do think we should color code it so we can talk about the individual parts. There’s no way to talk about each part without something differentiating them like color.

## Slide 27: CNT Filter Placement
Timothy Norman

## Slide 28: Carbon Nanotubes (CNTs)
Tubes made of carbon that have incredible tensile strength and thermal conductivity due to bonds between carbon atoms
Multi-Walled Carbon Nanotubes (MWCNTs) are nested nanotubes that provide unique diameter variability
Can be used in the form of yarns, fabrics, bucky paper and resin infused bucky paper
Kencin Autry

## Slide 29: Bucky Paper CNT Film
Kencin Autry
Multi-Walled Carbon Nanotube (MWCNT) Film
Average bucky paper thickness of about 0.06mm
Randomized dispersion with individual strands entangling
Notes: Multi-Walled Carbon Nanotube (MWCNT) Film
Bucky paper thickness of about 0.06mm
Randomized dispersion with individual strands entangling

## Slide 30: “Bucky paper” Production
Powdered CNTs
Sonicator
Vacuum Aided Filtration
Final Bucky Paper
Timothy Norman

## Slide 31: Bucky Paper Production
Sonication
Vacuum Filtration
Drying CNT film
Timothy Norman

## Slide 32: First Bucky Paper Production
Timothy Norman
Materials
Process Specifications
200 grams Isopropyl Alcohol
0.125 grams of MWCNTs
Vacuum filtered twice
Dried compressed in metal plates and Teflon paper
Sonicated for 3 hours and 15 minutes
Notes: 200 grams Isopropyl Alcohol
0.125 grams of powdered CNTs
Sonicated for a total of 3 hours and 15 minutes
Vacuum filtered twice
Vacuum aided drying overnight
Produced a sheet of thickness 0.03 millimeters
Dried with metal plates and Teflon paper
Also will include second round on same slide or on second

## Slide 33: Second Bucky Paper Production
Timothy Norman
Materials
Process Specifications
200 grams Isopropyl Alcohol
0.250 grams of MWCNTs
Sonicated for 6 hours and magnetically mixed overnight
Vacuum filtered twice
Dried compressed in metal plates and Teflon paper
Notes: 200 grams Isopropyl Alcohol
0.125 grams of powdered CNTs
Sonicated for a total of 3 hours and 15 minutes
Vacuum filtered twice
Vacuum aided drying overnight
Produced a sheet of thickness 0.03 millimeters
Dried with metal plates and Teflon paper
Also will include second round on same slide or on second

## Slide 34: Results
No filtration capabilities
Produced 3 sheets of CNTs all with equal thickness
CNTs were hydrophobic
Due to nature of entangled CNTs
Lapadre Proctor
Notes: Produced 3 sheets of CNTs all with equal thickness
The plan was to have three grates inside the housing and comparatively test better filtration with placing one by one inside
But..
Explain why we thought it failed
CNTs were hydrophobic
No filtration capabilities
Nature of randomly dispersed CNTs

## Slide 35: Response
Matthew Kennedy
Notes: Ordered a vertically aligned sheet of CNTs
Vertically aligned CNTs are a proven filtration method
Awaiting shipment

## Slide 36: Results
Also hydrophobic so no filtration capabilities
Used a pipette to drip water onto single film
Sheet too thin for desired flowrate
Due to bonds between the carbon atoms
Lapadre Proctor
Notes: CNT matrices are whats used in NASA prototype so might need to look at other forms of containing CNTs

## Slide 37: Testing Methods
TDS : 0 – 19.99 ppt
Conductivity
Water Temp
pH
Filtered Water
Matthew Kennedy
Notes: https://www.grainger.com/product/EXTECH-Conductivity-Meter-38W993

## Slide 38: Testing Methods
Filtered Water
Matthew Kennedy
Notes: Urea raises the pH and

## Slide 39: Further Testing Methods
Filtered Water
Total Organic Carbons
Resistivity
Ammonia
Matthew Kennedy
Notes: Total Organic Carbons
Reactivity
Endotoxins
Chemicals
Metals
$ 174 one week turnaround

## Slide 40: Testing Results
Before Reverse Osmosis
After Reverse Osmosis
TDS: 206 ppm
Conductivity: 322 μS/cm
Lead: ≈ 10 ppm
Carbonate: ≈ 240 ppm
pH: 7.6
TDS: 11.5 ppm
Conductivity: 18 μS/cm
Lead: ≈ 0 ppm
Carbonate: ≈ 0 ppm
pH: 6.0
Matthew Kennedy
Notes: Possible discrepancies in TDS and conductivity from the pen

## Slide 41: Testing Results
Spent Dialysate Model
After Spent Dialysate Through RO
TDS: 822 ppm
Conductivity: 1642 μS/cm
Lead: ≈ 0 ppm
Carbonate: ≈ 40 ppm
pH: 6.4
TDS: 59 ppm
Conductivity: 114 μS/cm
Lead: ≈ 0 ppm
Carbonate: ≈ 0 ppm
pH: 6.0
Matthew Kennedy

## Slide 42: Validated Targets
Raise and lower flow rate from 100-1000 mL/min
Recycling spent dialysate at 75% efficiency
Stores excess liquids and waste
Matthew Kennedy
Interface with existing ports
Notes: I like the animation! Recycling Spent Dialysate at 75% efficiency

## Slide 43: Unvalidated Targets
Tap water purifies to ultrapure standards through RO
System fitting within
10 x 12 x 5 inch dimensions
Matthew Kennedy

## Slide 44: Bill Of Materials
Chart: Updated Budget Of $2,000
 | 448.74 | 1551.26
Matthew Kennedy

## Slide 45: Bill Of Materials
Chart: Budget Consumption of $1,551.26
 | 473.59 | 1077.67
Matthew Kennedy

## Slide 46: Lessons Learned
Matthew Kennedy
Notes: Further research on deciding parts for the desk
Things going wrong that seemed unpredictable for the upside-down four-leaf clover
Not sure how to word this but sit down and do the math way earlier lol cause then we could have conducted experiments to find the unknown variables

## Slide 47
47
Future Solutions
Concept
Selection
Matthew Kennedy
CNT Matrix Membrane
Use More Filtration Steps
Use NASA Prototype
Machine CNT Housing
Notes: Not sure if this is the kind of things we wanted to put here

## Slide 48
Thank You
LinkedIn
Kencin Autry
Matt Kennedy
Lily Thompson
Lapadre Proctor
Timothy Norman
Notes: Add groups linked in info in QR code

## Slide 49: Backup Slides

## Slide 50: Targets and Metrics Table

## Slide 51: Ultra-Pure Standards Table

## Slide 52: Contaminants Found in Tap Water

## Slide 53
A mixed matrix (MM) CNTs membrane, on the other hand, consists of several layers of polymers or other composite material. Such membranes operate with low energy consumption due to the frictionless capacity of CNT to move water through hydrophobic hollow cavity nanotubes. The membrane is highly sensitive to multiple pollutants and salts, being permeable to antifouling, self-cleaning, reusable and good water. A brief comparison between the vertically aligned and mixed matrix CNT membranes is tabulated in Table 1.

## Slide 54: Standards for Ultra-Pure Water
Ultrapure water (UPW) is water that has been purified to high levels of specification:
  Extremely low total dissolved solids (TDS) level
  High electrical resistivity
  Free from chloride, sulfate, sodium, calcium, magnesium
  Minimal levels of organic contaminants, > 1 ppb

## Slide 55: Bill Of Materials
Matthew Kennedy
Chart: 
| Filter System | Pump System | Unspent | Testing
Budget Consumption | 1077.67 | 473.59 | 274.74 | 174

## Slide 56: Testing Methods
0.6 kilograms per meter cubed
Urea
0.45 gram per liter
Electrolyte Additives

## Slide 57: Nasa CNT and Ultrasonic Patent

## Slide 58: Initial vs Current Filter Prototype
CNT Membrane
Ultrasonic Generator
Surrounding Shell
Matthew Kennedy

## Slide 60: Fluid Computations
Lily Thompson
?
Key
?
Notes: Put dimensions on here

## Slide 61: Configuration
Lily Thompson
?
?
Notes: Put dimensions on here

## Slide 62: Concept Selection
Concept | Alternative Value
Nano-filtration and Reverse Osmosis | 0.247
Nano-filtration | 0.382
Forward Osmosis | 0.382
Analytical Hierarchy Process
Matt Kennedy

## Slide 63: Functional Decomp Priority Rating

## Slide 64: Functional Decomp Cross Reference

## Slide 65: Calculations for Laminar Flow
=
To find the velocity:
Given, by our customer, a flow rate of
1000 converted to
converted to
, indicating a Laminar Flow

## Slide 66: What a Dialyzer Looks Like

## Slide 69: Hemodialysis and Peritoneal
Hemodialysis is ongoing dialysis (3 to 5 times a week) that cleans your blood, usually in a dialysis center. The hemodialysis access is in your arm.
Peritoneal dialysis is ongoing dialysis (daily) that collects waste from the blood by washing the empty space in the abdomen (peritoneal cavity). It can be done from home. The access for this is in your abdomen.

## Slide 70: Reasons Why
The spent dialysate is a highly saline liquid consisting of urea, creatinine, uric acid, phosphate, etc.
A major challenge for the reuse of spent dialysate is removal of urea, the main nitrogenous waste product.
Due to its small molecular weight (60 Da) and electroneutral nature, it is difficult to be removed by sorbents.
Key Drawbacks:
System complexity
High energy cost
Membrane fouling and scaling

## Slide 71: Current System Schematic
Kencin Autry
Dialysis Machine
Water Supply
Ultra-Pure Runoff
Dialysate Runoff
Filter
Notes: *include specific parts we’ve ordered
Water supply will need to include the RO system
Filter will need to to be the CNT membrane
Pumps and piping being used

## Slide 72
Timothy Norman
Hospitals and Medical Care Facilities
At-home patients
72
Primary Markets

## Slide 73: Results
Produced 3 sheets of CNTs all with equal thickness
CNTs were hydrophobic
No filtration capabilities
Due to nature of entangled CNTs
Notes: Produced 3 sheets of CNTs all with equal thickness
The plan was to have three grates inside the housing and comparatively test better filtration with placing one by one inside
But..
Explain why we thought it failed
CNTs were hydrophobic
No filtration capabilities
Nature of randomly dispersed CNTs

## Slide 74
Nursing and Retirement Care
Pharmaceuticals
City Water Treatment
Food Manufacturing
74
Secondary Markets
Timothy Norman

## Slide 77
Notes: Filter process design

## Slide 79
Lily Thompson
79
Assumptions
Hemodialysis and peritoneal will be the main medical processes the system is used for
The system will only be used by one patient at a time
The incoming flow is laminar

## Slide 80
Lily Thompson
80
Assumptions
Hemodialysis and peritoneal will be the main medical processes the system is used for
The system will only be used by one patient at a time
The incoming flow is laminar

## Slide 81: College of Engineering Color Palette
Fang
Orange
RGB: 238, 118, 36
Hex: #EE7624
CMYK:2, 66, 99, 0
1 2 3 4
1 2 3 4
White
Pantone: PMS 000C
RGB: 255, 255, 255
Hex: #FFFFFF
CMYK: 0, 0, 0, 0
1 2 3 4
Black
Pantone: Black C
RGB: 0, 0, 0
Hex: #000000
CMYK:0, 0, 0, 100
Garnet
Pantone: PMS 195 C
RGB: 120, 47, 64
Hex: #782F40
CMYK:19, 90, 50, 55
1 2 3 4
Notes: Avoid 75% Black in B135

## Slide 82: Accent Color Palette
1 2 3 4
Tardis Blue
RGB: 0, 59, 111
Hex: #003B6F
CMYK:
1 2 3 4
Turquoise
RGB: 64, 224, 208
Hex: #40E0D0
CMYK:
1 2 3 4
American Orange
RGB: 255, 139, 0
Hex: #FF8B00
CMYK:
Rubine Red
RGB: 206, 0, 88
Hex: #CE0058
CMYK: 0, 100, 43, 12
1 2 3 4
1 2 3 4
Asagi-iro
RGB: 72, 146, 155
Hex: #48929b
CMYK:
1 2 3 4
Gainsboro
RGB: 220, 220, 220
Hex: #DCDCDC
CMYK:
1 2 3 4
Corn
RGB: 251, 236, 93
Hex: #FBEC5D
CMYK:
1 2 3 4
Timberwolf
RGB: 219, 215, 210
Hex: #DBD7D2
CMYK:
Imperial
RGB: 104, 40, 96
Hex: #682860
CMYK:
1 2 3 4
1 2 3 4
Tardis Blue
RGB: 0, 59, 111
Hex: #003B6F
CMYK:
Rubine Red
RGB: 206, 0, 88
Hex: #CE0058
CMYK: 0, 100, 43, 12
1 2 3 4
1 2 3 4
Corn
RGB: 251, 236, 93
Hex: #FBEC5D
CMYK:
1 2 3 4
Timberwolf
RGB: 219, 215, 210
Hex: #DBD7D2
CMYK:
Notes: Remove American orange add hyperlink grey
Make corn the last for b135
All work in 134 imperial move later
Tardis doesn’t work in b136

## Slide 83
https://color.adobe.com/create/color-wheel
Analogous
F7AB19
D67F15
EE7624
D64615
F73119
Monochromatic
6E3610
F0A16C
EE7624
6E4931
BASB1C
Triad
A14508
4BED3B
EE7624
250CED
2010A1
Complementary
FF8C40
EE7624
0098A1
24E2ED
A1470C
Split Complementary
28A164
2FED8D
EE7624
0848A1
1871ED
Double Split Complementary
3BED93
EE7624
0C6AED
ED2F18
EDAC2F
Square
C7ED3B
EE7624
0CE1ED
A418ED
ED660C
Compound
87724A
EE7624
60EFCF
09BA61
BA7E09
Shades
6E3610
EE7624
FA7625
D46820
AD551A
Notes: These are the variants that work well with Fang Orange. The bolded items show well on the projectors.

## Slide 84
https://color.adobe.com/create/color-wheel
Analogous
85412D
8F3831
782f40
8F3176
792D85
Monochromatic
C44D69
8E626D
782F40
C48796
451B25
Triad
C43959
78743B
782F40
236178
43A2C4
Complementary
C46078
782F40
25C43D
2F783A
C43959
Split Complementary
93C460
577835
782F40
39C49D
297861
Double Split Complementary
5A783B
782F40
237860
6E2978
784435
Square
78683b
782F40
237830
293978
782337
Compound
DEBAAF
782F40
5D8555
70AB32
AB4D32
Shades
C44D69
782F40
853447
5E2532
38161E

## Slide 85
B135 Stage Right
A105
A105
B135 Stage Left
B135
B134, A105
B136

## Slide 91: Logos and icons

## Slide 92: Citations
blood dialysis by Andy Horvath from <a href="https://thenounproject.com/browse/icons/term/blood-dialysis/" target="_blank" title="blood dialysis Icons">Noun Project</a> (CC BY 3.0)
Kidney by Febrian Hidayat from <a href="https://thenounproject.com/browse/icons/term/kidney/" target="_blank" title="Kidney Icons">Noun Project</a> (CC BY 3.0)
Dialysis Machine by bsd studio from <a href="https://thenounproject.com/browse/icons/term/dialysis-machine/" target="_blank" title="Dialysis Machine Icons">Noun Project</a> (CC BY 3.0)
Dialysis Machine by Gan Khoon Lay from <a href="https://thenounproject.com/browse/icons/term/dialysis-machine/" target="_blank" title="Dialysis Machine Icons">Noun Project</a> (CC BY 3.0)
filter by Smalllike from <a href="https://thenounproject.com/browse/icons/term/filter/" target="_blank" title="filter Icons">Noun Project</a> (CC BY 3.0)
filter by Kiran Shastry from <a href="https://thenounproject.com/browse/icons/term/filter/" target="_blank" title="filter Icons">Noun Project</a> (CC BY 3.0)
filter by Kiran Shastry from <a href="https://thenounproject.com/browse/icons/term/filter/" target="_blank" title="filter Icons">Noun Project</a> (CC BY 3.0)
Waste by lastspark from <a href="https://thenounproject.com/browse/icons/term/waste/" target="_blank" title="Waste Icons">Noun Project</a> (CC BY 3.0)
Waste by lastspark from <a href="https://thenounproject.com/browse/icons/term/waste/" target="_blank" title="Waste Icons">Noun Project</a> (CC BY 3.0)
Adapter by mynamepong from <a href="https://thenounproject.com/browse/icons/term/adapter/" target="_blank" title="Adapter Icons">Noun Project</a> (CC BY 3.0)
Outlet by Darrin Loeliger from <a href="https://thenounproject.com/browse/icons/term/outlet/" target="_blank" title="Outlet Icons">Noun Project</a> (CC BY 3.0)
Water Filter by I Putu Dicky Adi Pranatha from <a href="https://thenounproject.com/browse/icons/term/water-filter/" target="_blank" title="Water Filter Icons">Noun Project</a> (CC BY 3.0)
Water Filter by Phonlaphat Thongsriphong from <a href="https://thenounproject.com/browse/icons/term/water-filter/" target="_blank" title="Water Filter Icons">Noun Project</a> (CC BY 3.0)
water treatment by Soremba from <a href="https://thenounproject.com/browse/icons/term/water-treatment/" target="_blank" title="water treatment Icons">Noun Project</a> (CC BY 3.0)
water treatment by Soremba from <a href="https://thenounproject.com/browse/icons/term/water-treatment/" target="_blank" title="water treatment Icons">Noun Project</a> (CC BY 3.0)
Faucet by Uswa KDT from <a href="https://thenounproject.com/browse/icons/term/faucet/" target="_blank" title="Faucet Icons">Noun Project</a> (CC BY 3.0)
IV Bag by Ralf Schmitzer from <a href="https://thenounproject.com/browse/icons/term/iv-bag/" target="_blank" title="IV Bag Icons">Noun Project</a> (CC BY 3.0)
Reservoir by Adrien Coquet from <a href="https://thenounproject.com/browse/icons/term/reservoir/" target="_blank" title="Reservoir Icons">Noun Project</a> (CC BY 3.0)

## Slide 93: Citations
R. Nystrand Microbiology of water and fluids for hemodialysis J. Chin. Med. Assoc., 71 (5) (2008), pp. 223-229
J.W.M. Agar Conserving water in and applying solar power to haemodialysis: ’Green Dialysis’ through wiser resource utilization Nephrology, 15 (4) (2010), pp. 448-453
https://www.umms.org/ummc/health-services/kidney/dialysis/types#:~:text=Hemodialysis%20is%20ongoing%20dialysis%20(3,can%20be%20done%20from%20home.